REVIEW OF CRYSTAL CALORIMETERS
V.B.Golubev,
Budker Institute of Nuclear Physics, Novosibirsk, Russia

---
Crystal Ball Detector
The first large-scale crystal calorimeter in high energy physics was the NaI(Tl) calorimeter of the Crystal Ball detector. It showed high discovery potential of the crystal calorimeters.

---
Crystal Ball NaI(Tl) Calorimeter
Number of crystals 672
Inner radius 25.4 cm
Outer radius 66.0 cm
Thickness 16 X0
Solid angle coverage 93%
Photodetector PMT
Noise 0.05 MeV
Dynamic range 104
Inclusive photon spectrum at (2S) resonance

---
Crystal calorimeter benefits
High efficiencies for electrons and photons
Best energy resolution and good position and angular resolution for photons and electrons
High linearity and wide dynamic range
Radiation hardness
Reliability and simplicity of maintenance
Compact size

---
Energy resolution and line shape
Light collection and crystal light output nonuniformities
Fluctuations of the shower energy losses
Photoelectron statistics
Inaccuracy of the calibration of individual channels
Electronics noise and instabilities
Commonly used parameterizations of the calorimeter line shape:
BABAR calorimeter
line shape.
Fit – Crystal Ball function
N – norm. factor
x0 – peak position
–Gaussian width
– joint parameter
Crystal Ball
function
Logarithmic Gaussian
(Novosibirsk function)
N – norm. factor
x0 – peak position
–FWHM/2.35
– peak asymmetry

---
HEP Crystal Calorimeters
95–20
94–18
94–08
80–08
96–00
80–00
90–15
75–85
Time
104
-
PMT
27
2
-
3260
CsI
KTeV
Tevatron
Fermilab
105
0.8
Si PD
22
1.5
0.55
11400
BGO
L3, LEP, CERN
30
0.2
0.15
0.5
0.2
0.05
Noise (MeV)
105
APD, VPT
26
11
1.29
76000
PbWO4
CMS, LHC, CERN
104
104
104
103
104
Dynamic range
Si PD
Si PD
Si PD
VPT
PMT
Photo sensor
16.2
16-17.5
16
13.4
16
Thickness (X0)
9.5
5.9
7
1
1
Volume (m3)
1.25
0.9
1.0
0.25
0.254
Inner radius (m)
8800
6580
7800
1640
672
N crystals
CsI(Tl)
CsI(Tl)
CsI(Tl)
NaI(Tl)
NaI(Tl)
Crystal
BELLE KEK B KEK
BABAR PEP II SLAC
CLEO c, CESR, Cornell
SND,
VEPP-2M, VEPP-2000 BINP Novosibirsk
Crystal Ball,
SPEAR, SLAC
Experiment

---
Scintillation crystals for calorimetry
0.14
30
7.3
300
2.41
1.70
1.62
6.26
CeF3
1280
–2
~0
630
0.9
36
3.4
300
220
3.10
2.03
1.50
4.89
BaF2
-0.1
~0
-1.9
-1.6
-0.6
0.3
~0
dLY/dT
%/oC
1123
30
6
0.29
0.083
425
420
2.00
0.89
2.20
8.3
PWO
 (PbWO4)
2050
42
84
420
2.07
1.14
1.82
7.40
LSO
(Lu2SiO5:Ce)
1950
60
30
440
2.23
1.38
1.85
6.71
GSO
(Gd2SiO5:Ce)
1050
300
21
480
2.23
1.12
2.15
7.13
BGO
(Bi3Ge3O12)
621
30
6
3.6
1.1
420
310
3.57
1.86
1.95
4.51
CsI
621
1250
165
560
3.57
1.86
1.79
4.51
CsI(Tl)
230
Decay time
ns
100
Light yield
%NaI(Tl)
410
max
nm
4.13
Moliere
Radius
cm
2.59
Radiation
Length
cm
1.85
Refractive index
3.67
Density
g/cm3
Melting point
oC
Crystal
651
NaI(Tl)

---
SND detector
SND detector; 1 – Be vacuum chamber, 2 – tracking system, 3 – aerogel Cherenkov counters, 4 – NaI(Tl) scintillation counters, 5 – vacuum phototriodes, 6 – iron absorber, 7,8 – muon system.

---
SND NaI(Tl) calorimeter
SND calorimeter segment
SND calorimeter geometry
Number of crystals - 1632, thickness – 13.5 X0, mass – 3.5 tons
Photodetectors – vacuum phototriodes with photocathode quantum efficiency of ~15% and gain ~10.
Three-layer structure allows high quality e/π separation by using measurement of the shower energy deposition longitudinal profile

---
SND calorimeter resolution
Photon energy resolution
Angular resolution for photons
Crystal calibration – using cosmic muons (precalibration) and Bhabha scattering events (final calibration)
Future plans – the use of the calorimeter for the time of flight measurement for selection of e+e− ->nnbar events. The resolution of 2ns for energy deposition of 70 MeV was achieved in the test with an SND calorimeter crystal (to be presented at INSTR08 poster session)

---
L3 BGO calorimeter
Number of crystals:
    Barrel – 7680
    Endcaps – 2×1527
Thickness – 22X0
Photodotector – Si PD
L3 detector at LEP 1980-2000

---
L3 calorimeter calibration
Proton RFQ accelerator-based calibration system
RFQ accelerator produces 1.85 MeV H− beam. After focusing and steering it is neutralized and directed through the detector magnetic field to Li target inside the detector.
Resonant proton radiative capture reaction (σ=5mb) produces calibration photons:
p+7Li → 8Be + (17.6 MeV)
Calibration accuracy – 0.5%

---
L3 calorimeter energy resolution
0.88%
0.4%
0.5%
0.6%
Endcap
1.07%
0.5%
0.5%
0.8%
Barrel
Overall
Calibration
Temperature
Intrinsic
Contribution

---
L3 calorimeter discovery potential
Study of the process
e+e−→
Determination of the number
of light neutrinos

---
KTeV detector

---
KTeV CsI calorimeter
Calorimeter size 1.9x1.9 m2
Number of crystals 3100
Thickness 27 X0
Crystal size 2.5x2.5x50 cm3 (central part), 5.0x5.0x50 cm3 outer part
Photodetector – PMT
Calibration – electrons from Ke3 decays
Nonlinearity (3 – 75 GeV) 0.4%
Spacial resolution for photons ~1 mm

---
CLEO-c detector

---
CLEO-c CsI(Tl) calorimeter
Number of crystals 7784:
Barrel – 6144
Endcaps – 2×820
Crystal size – 5x5x30 cm3
Thickness – 16.2 X0
Inner radius – 1.02 m
Barrel length – 3.26 m
Total CsI(Tl) mass – 27000 kg
Solid angle coverage – 95%
Photodetector Si photodiode (4 per crystal)
Material in front of the barrel	part – 0.18X0

---
CLEO calorimeter performance
 mass resolution
Angular and energy resolution

---
BABAR Detector

---
BABAR CsI(Tl) Calorimeter
Number of crystals – 6580,
Thickness – 16÷17.5 X0
Inner radius – 0.92 m
Length – 2.9 m
Solid angle coverage 90% (in cms)

---
BABAR Calorimeter Calibration
Liquid radioactive source based on the reaction: 19F(n,)16N; 16N (7s)  16O*  16O+γ(6.13 MeV);
Bhabha scatterig (2.5 – 8 GeV), 200 direct hits per crystal, 0.35% statistical and 1% systematic error;
Additional corrections using π0 decays, e+e−→ e+e−γandµ+µ−γreactions
Neutron generator based on deuterium-tritium fusion reaction.
Generates up to 109 14 MeV n/s, when high voltage on.
Activates Fluorinert FC77 liquid
Counting rate 40 photons per crystal per second
Calibration accuracy ~0.25% for 15-30 minute run

---
BABAR Calorimeter Performance
peak in γγ invariant mass spectrum for E >300MeV
Liquid source calibration spectrum in one crystal.
Fit – 3 Gaussians (6.13 MeV peak + 2 escape peaks) and parameterized background.
Resolution – σ=0.3 MeV
Relative drop in light yield versus time

---
CMS Detector

---
CMS calorimeter requirements
One of the main goals of the CMS detector is a discovery of the Higgs boson. For the MH= 100÷150GeV the golden decay mode is H→.
The CMS EM calorimeter design resolution:
Radiation doses at high LHC luminosities:
Barrel center – 0.15Gy/h, Endcap – 15Gy/h
Typical loss of transparency in PbWO4 is 3% at 0.15 Gy/s – requires precision real time monitoring of the crystal optical properties.
Expected total -radiation doses for 10-year running at highest luminosity are 104 Gy in the barrel and ~106 Gy in endcaps. Expected neutron fluence is 1013 n/cm2

---
New crystals for HEP calorimetry
and ILC calorimeter proposal
R.Y. Zhu, 2005 ALCPG & ILC
Workshops Snowmass, USA[1]
LSO – Lu2SiO5 (Ce)
LYSO – Lu2(1-x)Y2xSiO5(Ce)
x=5 – 10%
In a last decade the mass production capabilities for LSO and LYSO crystals were established mostly for the medical industry. Their potential applications in HEP were also explored

---
LSO properties
Density – 7.40 g/cm3
Radiation length – 1.14 cm
Moliere radius – 2.07 cm
Light yield – 84% of NaI(Tl)
max – 420 nm
Decay time – 42 ns
Small temperature dependence of the light yield
0.511 MeV 22Na spectra from long crystal samples
Photodetector – 2 Hamamatsu 5x5 mm2 APDs per crystal.
Light output – 2000 p.e./MeV
Noise – 35 keV

---
LSO (LYSO) radiation hardness
The tests of radiation hardness of L(Y)SO crystals (J.M.Cheng et al., IEEE Trans Nucl Sci. 52 (2002)[2]; P.Kozma and P.Kosma Jr., NIM A539(2005)132[3]) showed that the radiation hardness of LSO crystals is very high.
     Complete recovery after 105Gy gamma irradiation takes few days. The radiation hardness of LSO is higher than that of BGO and PWO.
     The main problem with the LSO crystal radiation damage could be radiation-induced phosphorescence. The estimated radiation-induced read-out noise is ~1 MeV for 100 ns gate and 500 rad/h dose rate [2].
     Thus, the LSO and LYSO crystals are good candidates for ILC detector calorimeter. Another possible applicaion is the SuperB datector calorimeter endcaps.

---
Summary
In crystal calorimeters the e/ showers are totally absorbed in sensitive media. This provides the best energy resolution for electrons and photons and also a good spatial resolution.
Crystal calorimeters have wide dynamic range of up to 105 and low energy threshold of few MeV or less.
Crystal calorimeters are highly reliable and simple in maintenance.
New high-density scintillation crystals like LSO and LYSO have high light output and transparency, short scintillation decay time and very high radiation hardness. This makes the crystal calorometers based on these crystals a good choice for future HEP detectors including detector for ILC.